## Slide 1
Minerva Plant Evaluation
September 21st – October 1st 2015
Evaluation – September 2015

## Slide 2
Demonstration Objectives
  Concept Design
    Confirm features to measure
    Confirm visual fixture scheme
      Reference GD&T
      Reference Plant CMM fixture scheme
    Confirm all parts can be measured in plant environment
      Black/Gray/Blue porcelain
      Metal
  Physical Setup
    Determine where to locate system
    Setup area scanners
    Conduct Risk Assessment – Sundeep Jaganath
  System Validation
    Drift Check Results
    CMM Correlation Results
    Variation Reduction Analysis
  Next Steps
    General Notes
    Proposed Solution
Objectives

## Slide 3
Details
Drift Check
Error Proofing
  3 Measurement Locations
Model Measurement
  30” Minerva Model
    21 Checkpoints
      Datum A (3)
      Datum B (2)
      Datum C (1)
    Rail Width Positions
      Rail 2 Left/Right Front (2)
      Rail 2 Left/Right Middle (2)
      Rail 2 Left/Right Rear (2)
      Rail 5 Left/Right Front (2)
      Rail 5 Left/Right Middle (2)
      Rail 5 Left/Right Rear (2)
      Rail 7 Left/Right Front (2)
      Rail 7 Left/Right Middle (2)
      Rail 7 Left/Right Rear (2)
    9 Relationships
      Cavity Widths
        Front/Middle/Rear @ three rail positions
Concept Design: Overview

## Slide 4
Concept Design: GD&T Markup
General Capability Notes
The Cavity widths are measured at locations near the C1 through C9 Callouts.
Notes: Supplied By Chad Lako

## Slide 5
Concept Design: Features
A2
A1
B2
B1
A3
Whirlpool Datums
Datum A1
Datum A2
Datum A3
Datum B1
Datum B2
Datum C1
The Visual Fixture Scheme
has been confirmed to match that of the Minerva CMM.
C1

## Slide 6
L1/L4/L7
R1/R4/R7
C1/C4/C7
C2/C5/C8
C3/C6/C9
L2/L5/L8
R2/R5/R8
R3/R6/R9
L3/L6/L9
C1
C4
C7
C2
C5
C8
C3
C6
C9
X
Y
Z
Y
L1
L4
L7
R1
R4
R7
l
l
l
l
l
l
C1: distance between L1 and R1
C4: distance between L4 and R4
C7: distance between L7 and R7
l
l
l
l
l
l
l
l
l
l
l
l
L8
L5
L2
L9
L6
L3
R8
R5
R2
R9
R6
R3
Concept Design: Measurement Locations

## Slide 7
Physical Setup
Physical Setup
  Fixed Locators handle all Parts
  Area Scanner Volume Adjusted for Measurement Area
  Risk Assessment Comducted – Sept. 22nd (Sundeep Jaganath)
  Robot Paths & Pendant Display Tested
    Fanuc M-10iA 12 Robot
      DCS Robot
    Cycle Times
      Fiducial Measurement ~ 8 seconds
      Error Proofing ~ 14 seconds
      Part Measurement ~ 102 seconds
    Robot Cycling Speed at 50%
    of Robot Maximum Speed due to Risk Assessment Report

## Slide 8
Video: Full Routine – Blue Cavity

## Slide 9
System Validation – Drift Check Results
Results
  233 Drift Check Cycles from 9/22 to 9/29
  Range of .17 mm
  A couple flyers when system starts up
  Initial analysis is Robot Temperature Compensation is not needed.

## Slide 10
System Validation – Correlation Study
Results
  10 Parts run through Perceptron Station and CMM
  All 3 Porcelain Colors were Measured
  Correlation is between width distance across cavity opening
    2 CMM feature locations
    2 Perceptron surface algorithms
  Worst Correlation Coefficient at 0.956
  Maximum Offset of only 0.128
  285 Cavities have been measured as of 1:33PM 9/30.
Perc Axis | CMM-Axis | Perc Mean | CMM Mean | Correlation Coefficient | 6 Sigma | Calculated Offset
C1 | D1 | 632.628 | 632.662 | 0.999 | 0.442 | 0.034
C2 | D2 | 633.076 | 632.971 | 1.000 | 0.433 | -0.105
C3 | D3 | 633.059 | 633.112 | 0.997 | 0.421 | 0.053
C4 | D4 | 633.361 | 633.430 | 0.998 | 0.323 | 0.069
C5 | D5 | 633.872 | 633.897 | 0.999 | 0.583 | 0.025
C6 | D6 | 633.665 | 633.735 | 0.985 | 0.271 | 0.070
C7 | D7 | 633.657 | 633.639 | 0.999 | 0.559 | -0.018
C8 | D8 | 633.654 | 633.782 | 0.996 | 0.343 | 0.128
C9 | D9 | 634.084 | 634.177 | 0.956 | 0.375 | 0.093
C1

## Slide 11
System Validation – Correlation Study
Results
  10 Parts run through Perceptron Station and CMM
  All 3 Porcelain Colors were Measured
  Correlation is between width distance across cavity opening
    2 CMM feature locations
    2 Perceptron surface algorithms
  Worst Correlation Coefficient at 0.956
  Maximum Offset of only 0.128
Perc Axis | CMM-Axis | Perc Mean | CMM Mean | Correlation Coefficient | 6 Sigma | Calculated Offset
C1 | D1 | 632.628 | 632.662 | 0.999 | 0.442 | 0.034
C2 | D2 | 633.076 | 632.971 | 1.000 | 0.433 | -0.105
C3 | D3 | 633.059 | 633.112 | 0.997 | 0.421 | 0.053
C4 | D4 | 633.361 | 633.430 | 0.998 | 0.323 | 0.069
C5 | D5 | 633.872 | 633.897 | 0.999 | 0.583 | 0.025
C6 | D6 | 633.665 | 633.735 | 0.985 | 0.271 | 0.070
C7 | D7 | 633.657 | 633.639 | 0.999 | 0.559 | -0.018
C8 | D8 | 633.654 | 633.782 | 0.996 | 0.343 | 0.128
C9 | D9 | 634.084 | 634.177 | 0.956 | 0.375 | 0.093
C2

## Slide 12
System Validation – Correlation Study
Results
  10 Parts run through Perceptron Station and CMM
  All 3 Porcelain Colors were Measured
  Correlation is between width distance across cavity opening
    2 CMM feature locations
    2 Perceptron surface algorithms
  Worst Correlation Coefficient at 0.956
  Maximum Offset of only 0.128
Perc Axis | CMM-Axis | Perc Mean | CMM Mean | Correlation Coefficient | 6 Sigma | Calculated Offset
C1 | D1 | 632.628 | 632.662 | 0.999 | 0.442 | 0.034
C2 | D2 | 633.076 | 632.971 | 1.000 | 0.433 | -0.105
C3 | D3 | 633.059 | 633.112 | 0.997 | 0.421 | 0.053
C4 | D4 | 633.361 | 633.430 | 0.998 | 0.323 | 0.069
C5 | D5 | 633.872 | 633.897 | 0.999 | 0.583 | 0.025
C6 | D6 | 633.665 | 633.735 | 0.985 | 0.271 | 0.070
C7 | D7 | 633.657 | 633.639 | 0.999 | 0.559 | -0.018
C8 | D8 | 633.654 | 633.782 | 0.996 | 0.343 | 0.128
C9 | D9 | 634.084 | 634.177 | 0.956 | 0.375 | 0.093
C3

## Slide 13
System Validation – Correlation Study
Results
  10 Parts run through Perceptron Station and CMM
  All 3 Porcelain Colors were Measured
  Correlation is between width distance across cavity opening
    2 CMM feature locations
    2 Perceptron surface algorithms
  Worst Correlation Coefficient at 0.956
  Maximum Offset of only 0.128
Perc Axis | CMM-Axis | Perc Mean | CMM Mean | Correlation Coefficient | 6 Sigma | Calculated Offset
C1 | D1 | 632.628 | 632.662 | 0.999 | 0.442 | 0.034
C2 | D2 | 633.076 | 632.971 | 1.000 | 0.433 | -0.105
C3 | D3 | 633.059 | 633.112 | 0.997 | 0.421 | 0.053
C4 | D4 | 633.361 | 633.430 | 0.998 | 0.323 | 0.069
C5 | D5 | 633.872 | 633.897 | 0.999 | 0.583 | 0.025
C6 | D6 | 633.665 | 633.735 | 0.985 | 0.271 | 0.070
C7 | D7 | 633.657 | 633.639 | 0.999 | 0.559 | -0.018
C8 | D8 | 633.654 | 633.782 | 0.996 | 0.343 | 0.128
C9 | D9 | 634.084 | 634.177 | 0.956 | 0.375 | 0.093
C4

## Slide 14
System Validation – Correlation Study
Results
  10 Parts run through Perceptron Station and CMM
  All 3 Porcelain Colors were Measured
  Correlation is between width distance across cavity opening
    2 CMM feature locations
    2 Perceptron surface algorithms
  Worst Correlation Coefficient at 0.956
  Maximum Offset of only 0.128
Perc Axis | CMM-Axis | Perc Mean | CMM Mean | Correlation Coefficient | 6 Sigma | Calculated Offset
C1 | D1 | 632.628 | 632.662 | 0.999 | 0.442 | 0.034
C2 | D2 | 633.076 | 632.971 | 1.000 | 0.433 | -0.105
C3 | D3 | 633.059 | 633.112 | 0.997 | 0.421 | 0.053
C4 | D4 | 633.361 | 633.430 | 0.998 | 0.323 | 0.069
C5 | D5 | 633.872 | 633.897 | 0.999 | 0.583 | 0.025
C6 | D6 | 633.665 | 633.735 | 0.985 | 0.271 | 0.070
C7 | D7 | 633.657 | 633.639 | 0.999 | 0.559 | -0.018
C8 | D8 | 633.654 | 633.782 | 0.996 | 0.343 | 0.128
C9 | D9 | 634.084 | 634.177 | 0.956 | 0.375 | 0.093
C5

## Slide 15
System Validation – Correlation Study
Results
  10 Parts run through Perceptron Station and CMM
  All 3 Porcelain Colors were Measured
  Correlation is between width distance across cavity opening
    2 CMM feature locations
    2 Perceptron surface algorithms
  Worst Correlation Coefficient at 0.956
  Maximum Offset of only 0.128
Perc Axis | CMM-Axis | Perc Mean | CMM Mean | Correlation Coefficient | 6 Sigma | Calculated Offset
C1 | D1 | 632.628 | 632.662 | 0.999 | 0.442 | 0.034
C2 | D2 | 633.076 | 632.971 | 1.000 | 0.433 | -0.105
C3 | D3 | 633.059 | 633.112 | 0.997 | 0.421 | 0.053
C4 | D4 | 633.361 | 633.430 | 0.998 | 0.323 | 0.069
C5 | D5 | 633.872 | 633.897 | 0.999 | 0.583 | 0.025
C6 | D6 | 633.665 | 633.735 | 0.985 | 0.271 | 0.070
C7 | D7 | 633.657 | 633.639 | 0.999 | 0.559 | -0.018
C8 | D8 | 633.654 | 633.782 | 0.996 | 0.343 | 0.128
C9 | D9 | 634.084 | 634.177 | 0.956 | 0.375 | 0.093
C6

## Slide 16
System Validation – Correlation Study
Results
  10 Parts run through Perceptron Station and CMM
  All 3 Porcelain Colors were Measured
  Correlation is between width distance across cavity opening
    2 CMM feature locations
    2 Perceptron surface algorithms
  Worst Correlation Coefficient at 0.956
  Maximum Offset of only 0.128
Perc Axis | CMM-Axis | Perc Mean | CMM Mean | Correlation Coefficient | 6 Sigma | Calculated Offset
C1 | D1 | 632.628 | 632.662 | 0.999 | 0.442 | 0.034
C2 | D2 | 633.076 | 632.971 | 1.000 | 0.433 | -0.105
C3 | D3 | 633.059 | 633.112 | 0.997 | 0.421 | 0.053
C4 | D4 | 633.361 | 633.430 | 0.998 | 0.323 | 0.069
C5 | D5 | 633.872 | 633.897 | 0.999 | 0.583 | 0.025
C6 | D6 | 633.665 | 633.735 | 0.985 | 0.271 | 0.070
C7 | D7 | 633.657 | 633.639 | 0.999 | 0.559 | -0.018
C8 | D8 | 633.654 | 633.782 | 0.996 | 0.343 | 0.128
C9 | D9 | 634.084 | 634.177 | 0.956 | 0.375 | 0.093
C7

## Slide 17
System Validation – Correlation Study
Results
  10 Parts run through Perceptron Station and CMM
  All 3 Porcelain Colors were Measured
  Correlation is between width distance across cavity opening
    2 CMM feature locations
    2 Perceptron surface algorithms
  Worst Correlation Coefficient at 0.956
  Maximum Offset of only 0.128
Perc Axis | CMM-Axis | Perc Mean | CMM Mean | Correlation Coefficient | 6 Sigma | Calculated Offset
C1 | D1 | 632.628 | 632.662 | 0.999 | 0.442 | 0.034
C2 | D2 | 633.076 | 632.971 | 1.000 | 0.433 | -0.105
C3 | D3 | 633.059 | 633.112 | 0.997 | 0.421 | 0.053
C4 | D4 | 633.361 | 633.430 | 0.998 | 0.323 | 0.069
C5 | D5 | 633.872 | 633.897 | 0.999 | 0.583 | 0.025
C6 | D6 | 633.665 | 633.735 | 0.985 | 0.271 | 0.070
C7 | D7 | 633.657 | 633.639 | 0.999 | 0.559 | -0.018
C8 | D8 | 633.654 | 633.782 | 0.996 | 0.343 | 0.128
C9 | D9 | 634.084 | 634.177 | 0.956 | 0.375 | 0.093
C8

## Slide 18
System Validation – Correlation Study
Results
  10 Parts run through Perceptron Station and CMM
  All 3 Porcelain Colors were Measured
  Correlation is between width distance across cavity opening
    2 CMM feature locations
    2 Perceptron surface algorithms
  Worst Correlation Coefficient at 0.956
  Maximum Offset of only 0.128
Perc Axis | CMM-Axis | Perc Mean | CMM Mean | Correlation Coefficient | 6 Sigma | Calculated Offset
C1 | D1 | 632.628 | 632.662 | 0.999 | 0.442 | 0.034
C2 | D2 | 633.076 | 632.971 | 1.000 | 0.433 | -0.105
C3 | D3 | 633.059 | 633.112 | 0.997 | 0.421 | 0.053
C4 | D4 | 633.361 | 633.430 | 0.998 | 0.323 | 0.069
C5 | D5 | 633.872 | 633.897 | 0.999 | 0.583 | 0.025
C6 | D6 | 633.665 | 633.735 | 0.985 | 0.271 | 0.070
C7 | D7 | 633.657 | 633.639 | 0.999 | 0.559 | -0.018
C8 | D8 | 633.654 | 633.782 | 0.996 | 0.343 | 0.128
C9 | D9 | 634.084 | 634.177 | 0.956 | 0.375 | 0.093
C9

## Slide 19
System Validation – Measurement Monitor Display
Line Display – Cavity Widths

## Slide 20
Distance 6-sigma variation Pareto chart based on Data collected on 9/28/2015
C1
C4
C7
C2
C5
C8
C3
C6
C9
l
l
l
l
l
l
l
l
l
l
l
l
l
l
l
l
l
l
C8
C5
C2
C7
C4
C1
C9
C6
C3
System Validation – Variation Reduction Analysis

## Slide 21
Next Steps
General Notes:
    100% In-Line Measurement
    10 Second Cycle Time
    Model Mix Needs to be Identified
      Minerva, Urso, Scott Line
      Determines number of Required Stations
    Possibility of Individual Part Identification (Barcode Scanner) in Future
Proposed Solution:
    100% In-Line Automated Solution
    1 or 2 Stations depending on Model Mix
    Servo Slide Solution(s)
      4 Sensors
      3 Positions
    Located after Current Sensing Station
    Sensing Station Requirements
      Determine part on carrier
      Confirm parts are located correctly on carriers
      Information to be sent to PLC
    Repair Loop to be moved after last Perceptron Station
    Provide a list of Preferred Local Integrators